## Slide 1: Neuroergonomics of an At-Home VR-based Psychophysical Procedure Compared to Clinical Visual Field Perimetry
Ronald Zambrano, MSc, PhD(c)
6th International Neuroergonomics Conference — Boston, MA
Wills Eye Hospital, Glaucoma Service, Philadelphia, PA, USA
Drexel University, School of Biomedical Engineering and Science, Philadelphia, PA, USA
Session 1: Clinical Neuroergonomics & Brain Health

## Slide 2: Financial Disclosure
I have no financial interests to disclose

## Slide 3: The Visual Field Test
Perimetry, quantitatively maps functional vision
  Standard of care across ophthalmology
  Guides diagnosis, monitoring, treatment
  Critical for irreversible diseases like glaucoma
Informs where and how much a patient sees
Detects vision loss, even when patients can’t
Visual function reference standard for decades1
zeiss.com
1. Li, A., & De Moraes, C. G. (2025). Ophthalmology. Glaucoma, 8(5S), S20–S26.

## Slide 4
There is one problem…

## Slide 5: The Problem with Standard Perimetry
Cognitively demanding
  Constant vigilance + Immediate user input
  Ergonomic burden: chin rest, one eye covered, dimly lit, clinic-bound
  Fatigue (Duration:10-25 min)
High variability
  Up to 40% unreliable tests in some studies2
Human-factors problem: test design, not the patient, causes most of the issues
2. Birt, C. M. et al., (1997). Ophthalmology, 104(7), 1126–1130.
willseyehospital.org
Notes:   Sustained fixation + constant vigilance
  Unnatural task: suppress the urge to look at what you see
Cognitively demanding
  Sustained fixation + constant vigilance
  Repeated immediate user feedback
  Fatigue (Duration:10-25 min)
  Ergonomic burden: chin rest, one eye covered, clinic-bound
Highly variable
  Up to 40% unreliable tests in some studies2

## Slide 6: Vivid Vision Perimetry (VVP) A human-factors redesign
VR headset, performed anywhere
  Self-paced · rest allowed · both eyes
Uses natural fixation approach
  Localize, don't just detect
Adaptive: learns a priori VF defects to reduce future test duration
Perimetry redesigned to
  Meet current standard reference
  Decrease neurocognitive burden
seevividly.com
Notes:   Sustained fixation + constant vigilance
  Unnatural task: suppress the urge to look at what you see
Cognitively demanding
  Sustained fixation + constant vigilance
  Repeated immediate user feedback
  Fatigue (Duration:10-25 min)
  Ergonomic burden: chin rest, one eye covered, clinic-bound
Highly variable
  Up to 40% unreliable tests in some studies2

## Slide 7: Study Aim
Overall Aim: Evaluate Vivid Vision Perimetry (VVP) as an alternative to Standard Automated Perimetry (SAP) in a glaucoma cohort
Specific Aim 1: To determine whether VVP produces measurements that agree with SAP
Specific Aim 2: To assess whether VVP reduces neuroergonomic burden across symptom, usability, preference, and confidence

## Slide 8: Methods

## Slide 9: Study Design & Analysis
1
Enrollment
Prospective Study
Wills Eye + UCSF
Glaucoma/ Healthy
2a
VVP
In-clinic training
Bundle of 10 sessions
~14 days
~18 min each
2b
SAP
4× Humphrey 24-2 SITA (Zeiss)
Across 2 clinic visits within 14 days
3
Surveys
After training & after bundle
Symptoms
Usability
Preference
Confidence
Agreement
Pearson correlation
Analysis
Comparison
Wilcoxon signed-rank test

## Slide 10: Results

## Slide 11: Study Cohort
Characteristic | Value
Enrolled participants (Glaucoma/ Health) | 63 (52/11)
Age, mean (SD), y | 59.4 (14.8)
Women / Men, n | 41 (65%) / 22 (35%)
Site: Wills / UCSF | 37 / 26
Completed home bundle (≥10 sessions ≤14 d) | 54 (85.7%)
Provided post-bundle survey data | 55 (87.3%)
Did not complete bundle / lost to follow-up | 9 (14.3%)
Retention defined as completion of the 10-session home bundle within 14 days.

## Slide 12: VVP and SAP reveal similar VF loss
VVP
SAP​
Right Eye
Left Eye

## Slide 13: VVP agrees with SAP
Strong Correlation of Average mean sensitivity: 10 VVP vs 4 SAP
3. Shiu et al., (2026), IOVS Vol.67, 5572
Similar strong correlation is seen pointwise (r = 0.70 – 0.95)

## Slide 14: Patients reported lower symptom burden
Post-test symptom severity (0–4; lower is better) after at-home VVP vs in-clinic SAP
Every domain significantly favored VVP
Symptom | VVP | SAP | p
Eye burning | 0.45 | 0.79 | 0.008
Eye dryness | 0.84 | 1.29 | 0.003
Eye strain | 0.93 | 1.44 | 0.003
Eye tearing | 0.42 | 0.78 | 0.004
Eye tiredness | 1.22 | 1.71 | 0.004
Neck/back soreness | 0.84 | 1.35 | 0.034
Non-parametric Wilcoxon signed-rank test/ Benjamini-Hochberg-corrected
Values are mean (SD) symptom-severity scores on the study’s ordinal scale; lower scores indicate less symptom burden.

## Slide 15: Patients preferred VVP
After the home bundle, most patients preferred VVP to SAP;
  preference grew from training day to post-bundle
  Top drivers: time savings, ability to rest, both eyes at once, convenience
Patient Confidence rose for VVP (3.6→4.1, p<0.01)
Chart: VVP vs SAP Preference
| Training Day (n=56) | After Bundle 1 (n=55)
Strongly prefer VVP | 0.482 | 0.564
Slightly prefer VVP | 0.196 | 0.2
No preference | 0.232 | 0.109
Slightly prefer HVF | 0.071 | 0.127
Strongly prefer HVF | 0.018 | 0
* preference scored on a Likert scale (0-4)

## Slide 16: Discussion

## Slide 17: Discussion
At-home VVP agrees closely with the SAP, overall and point-by-point
Patients report less symptom burden, prefer VVP after real use, and trust its results more than standard clinical SAP
Limitations:
  Unequal test count and strategy: 10 VVP vs 4 SAP; Not true 1:1 comparison
  Surveys: study-specific; Likert/ordinal; small (significant) magnitude differences
  Modest sample: Only testing early experience. Will findings hold over time?

## Slide 18: Future Directions
Quantify neurocognitive effects on perimetry
Currently enrolling for a multimodal Perimetric-Neuroimaging study
  fNIRS monitoring neural correlates during Perimetry (SAP & VVP)
For more information, please visit our poster
  Session: Applications, Safety & Ethics in Real-world Domains [H4-Wednesday]

## Slide 19: Conclusion
Designing the perimetric task around human attention and natural eye movement reduces burden without sacrificing measurement quality
Supports VVP as a patient-centered alternative for remote VF monitoring
Moving a demanding clinical test into the patient’s everyday world! 🌎

## Slide 20: Thank You!
For more information:
rzambrano@willseye.org
willseye.org/acorn-lab
- R44-EY035203
- R01-EY013178
Wills Eye: Yoav Gutterman; Vivin Mathew; Heather Giusto; Ahna Donahue; Batsheva Rubin; Bo Young Kim; Gadi Wollstein; Fabio Lavinsky; Joel S. Schuman
Vivid Vision: Benjamin T. Backus; Thaddeus B. Czuba; James Blaha; Ranran L. French
UCSF: Corinne Shiu; Michael S. Deiner; Yvonne Ou
Drexel: Special thanks to Dr. Hasan Ayaz and Dr. Adrian Curtin for the advice and guidance!